## Slide 1: Use of BinaxNOW COVID-19 Antigen (Ag) Test Training for Long Term Care Facilities
Kentucky Department for Public Health and Abbott Laboratories
December 16, 2020
1:00 - 2:15 P.M. (EST)
Updated #38 on 1/120/21

## Slide 2: Key Changes: CHFS Surveillance Testing Program for Long-Term Care Facilities (12/14/20 Guidance Memo)
Discontinues After 12/30/2020
CARES Act Funding
Contract: PCR Sampling
Contract: PCR Assays
  PCR validation: Antigen-Positive
Also - Current CHFS Contractors:
 Execute Modified Contract
(Technical Corrections)
Begin Immediately
 (Lab) Execute NEW Contract
 (LTCF) Alternate PCR/POC
(thru 12/30)
CHFS Resources Begin 12/31/2020
 Medicare-Certified: BinaxNow™
Order Qty in Dec: Dec-Feb
 Non-Med-Certified:
PCR Bi-Weekly (Clinical Lab)
Keith Knapp, Office of the CHFS Secretary
Notes: K2: brief overview of 12/14/20 Guidance memo

## Slide 3: Please Review & Refer To:
12/14/20 CHFS Provider Guidance Memo:
LTC COVID-19 Surveillance Testing Program at: https://chfs.ky.gov/cv19/LTCTestingProgramProviderGuidanceUpdate.pdf
AND
BinaxNOW™ Resource Request Form for Long Term Care Facilities
Available thru 12/30/2020 at:
https://ky.readyop.com/fs/4iY5/ac9d
Notes: New Slide 12/18

## Slide 4: Abbott BinaxNOWTM COVID-19 Antigen (Ag) Cards
Michelle Moran, Technical Consultant, Abbott Rapid Diagnostics

## Slide 5: BinaxNOW™
COVID-19 Ag Card
Overview
Emergency Use Authorization Granted by FDA

## Slide 6: Intended Use: Key Points
Proprietary and confidential — do not distribute
The BinaxNOW™ COVID-19 Ag Card is intended for the qualitative detection of nucleocapsid protein antigen from SARS-CoV-2 in nasal swabs from individuals suspected of COVID-19 by their healthcare provider within the first seven days of symptom onset.
Antigen is generally detectable in nasal swabs during the acute phase of infection.
Negative results from patients with symptom onset beyond seven days, should be treated as presumptive and confirmation with a molecular assay if necessary, for patient management, may be performed.

## Slide 7: BinaxNOW™ COVID-19 Ag Card Emergency Use Authorization
Proprietary and confidential — do not distribute
The BinaxNOW™ COVID-19 Ag Card is only for use under the Food and Drug Administration’s Emergency Use Authorization
What is Emergency Use Authorization (EUA)?
FDA emergency access mechanism
Health & Human Services declare when circumstances exist to justify use of diagnostics under EUA for the diagnosis of COVID-19
It is not full FDA clearance or approval and is temporary, until the declaration is terminated or revoked.

## Slide 8: BinaxNOW™ COVID-19 Ag Card Emergency Use Authorization
Proprietary and confidential — do not distribute
Test Site Obligations:
Notify relevant public health authorities on intent to run test
Report all results to healthcare providers and include the Healthcare Provider Fact Sheet. Healthcare providers to include Patient Fact Sheet with results
Utilize product as outlined in the BinaxNOW™ COVID-19 Ag Card Instructions for Use
Ensure all operators are trained to perform and interpret the test
Have process in place for reporting test result to Healthcare Providers & relevant public health authorities
Per Product Insert: Collect performance data and report any significant deviations from the product performance characteristics via email to FDA/HHS and to Abbott Technical Support
Retain all records associated with EUA until otherwise directed by FDA

## Slide 9: BinaxNOW™ COVID-19 Ag Card Technical Overview

## Slide 10: BinaxNOW™ COVID-19 Ag Card Product Overview
Proprietary and confidential — do not distribute
Test Summary | Rapid lateral flow immunoassay for the qualitative detection of SARS-CoV-2
Testing Environment | Point of Care settings with a CLIA Certificate of Waiver, Compliance or Accreditation
Specimen Type | Direct nasal swab
Time to Result | Results visually read at 15 minutes *Results should not be read after 30 minutes
Waste Disposal | All components should be discarded as Biohazard Waste
PPE for Specimen Collection & Handling | Refer to CDC Guidelines for collecting, handling and testing clinical specimens (link in Product Insert)

## Slide 11: Reagent and Materials
Proprietary and confidential — do not distribute
Materials Provided:
40 Test Cards
Extraction Reagent
Patient Collection Nasal Swabs
Positive Control Swab
Blank Nasal Swab for Negative Control
Product Insert
Procedure Card
Healthcare Provider & Patient COVID-19 Fact Sheets
Materials Required but not Provided:
Clock, timer or stopwatch
Optional Materials:
Plastic Transport Tube
Storage & Stability:
Store kit at 2-30°C
Ensure all test components are at room temperature before use
Stable until the expiration date marked on the outer packaging

## Slide 12: Quality Control
Proprietary and confidential — do not distribute
Internal Procedural Controls:
BinaxNOW™ COVID-19 Ag Card has built-in procedural controls
In an untested card there will be a blue line at the Control Line position
In a valid, tested device, the blue line washes away and a pink/purple line appears, confirming that the sample has flowed through the test strip and the reagents are working
Note: If the blue line is not present at the Control Line position prior to running the test, do not use and discard

## Slide 13: When is Quality Control Required?
Proprietary and confidential — do not distribute
External Positive & Negative Controls:
Good laboratory practice suggests the use of positive and negative controls to ensure that test reagents are working properly and that the test is correctly performed
BinaxNOW™ COVID-19 Ag Card kits contain a positive control swab and sterile swabs that can be used as a negative control
Required Frequency:
New shipments received
Untrained operators
Conforming with local, state, and/or federal regulations, accrediting groups, or lab’s standard QC procedures.
Note: If correct results are not obtained, contact the Abbott Technical Services Team at
1-800-257-9525 between 8 a.m. and 8 p.m. EST Monday-Friday before testing patient specimens

## Slide 14: Nasal Swab Sample Collection
Insert the nasal swab into the nostril exhibiting the most drainage or congestion
Using gentle rotation, push the swab until resistance is met
  At the level of the nasal turbinates
  Less than one inch into nostril
Rotate the swab 5 times or more against the nasal wall
Slowly remove the swab
Using the same swab, repeat sample collection in the other nostril
Only the swab provided in the kit is to be used for nasal swab collection
Notes: Mention: Qualifications for performing swab
(No certification/credential required in KY; document proof of training)
Reference to a YouTube video?

## Slide 15: Specimen Transport & Storage

## Slide 16: Specimen Transport & Storage
For best performance, direct nasal swabs should be tested as soon as possible after collection.
If immediate testing is not possible:
Do not return the nasal swab to the original paper packaging
To avoid contamination and preserve sample integrity, place the nasal swab in a clean, unused, tightly capped plastic tube & label with the patient information
The sample is stable in the plastic tube at room temperature (15-30°C) for up to one (1) hour prior to testing
If greater than one (1) hour delay occurs, dispose of sample & collect new sample

## Slide 17: BinaxNOW™ COVID-19 Ag Card Procedure
Proprietary and confidential — do not distribute
Key Points
Prior to Beginning a Test Remember to:
Ensure all components are at room temperature
Open test card j ust prior to use
Lay test card flat for use
Ensure Blue Control Line is present

## Slide 18: BinaxNOW™ COVID-19 Ag Card Overview
Exterior View
Interior View

## Slide 19: BinaxNOW™ COVID-19 Ag Card Test Procedure Overview
Proprietary and confidential — do not distribute

## Slide 20: BinaxNOW™ COVID-19 Ag Card Procedure Key Points
Proprietary and confidential — do not distribute
Step 1:
Hold extraction reagent bottle vertically to ensure adequate volume
Add 6 drops for a Patient test
Add 8 drops for a Quality Control test

## Slide 21: BinaxNOW™ COVID-19 Ag Card Procedure Key Points
Proprietary and confidential — do not distribute
Step 2:
Insert swab into the bottom hole of the test card
Firmly push swab upwards until the swab tip is visible in the top hole

## Slide 22: BinaxNOW™ COVID-19 Ag Card Procedure Key Points
Proprietary and confidential — do not distribute
Step 3:
Rotate swab shaft 3 times CLOCKWISE (to the right).
False negative results can occur if the sample swab is not rotated (twirled) prior to closing the card

## Slide 23: BinaxNOW™ COVID-19 Ag Card Procedure Key Points
Proprietary and confidential — do not distribute
Step 4:
Peel off adhesive liner from the right edge of the test card
Close and securely seal the card
To ensure proper test performance read results at 15 minutes and not before
Results should not be read after 30 minutes

## Slide 24: Result Interpretation
Proprietary and confidential — do not distribute
Negative
One pink/purple colored line in the top half of the window, in the Control Line position
Indicates a negative result & no antigen detected
Positive
Two pink/purple colored lines in both the Control & Sample Line position
Indicates COVID-19 antigen was detected
Any visible pink/purple line is positive

## Slide 25: Result Interpretation
Proprietary and confidential — do not distribute
Invalid
The assay is invalid & should be repeated if:
Only the Sample Line is present
No lines are present
Blue Control Line remains blue
Test should be discarded as Biohazard waste

## Slide 26: Near-Patient Workflow: No Swab Transport*
Proprietary and confidential — do not distribute
Option A (Collect Patient Swab First, Prior to Test Procedure)
Step 1: Add the Extraction Reagent
Step 2: Insert sample into bottom hole and push firmly upwards so tip is visible in top hole
Step 3: Rotate (twirl) swab shaft 3 times clockwise (to the right)
Step 4: Peel off adhesive liner, close and securely seal card. Read results in the window 15 minutes after closing the card
Collect Nasal Swab
*Swab will not be moved from immediate testing area where sample collection is performed
Collect Nasal Swab

## Slide 27: Near-Patient Workflow: No Swab Transport* Option B (Begin Test Procedure, Then Collect Patient Swab)
Proprietary and confidential — do not distribute
Step 1: Add the Extraction Reagent
Step 2: Insert sample into bottom hole and push firmly upwards so tip is visible in top hole
Step 3: Rotate (twirl) swab shaft 3 times clockwise (to the right)
Step 4: Peel off adhesive liner, close and securely seal card. Read results in the window 15 minutes after closing the card
Collect Nasal Swab
Within 3 minutes of adding extraction reagent drops, insert swab into test card
*Swab will not be moved from immediate testing area where sample collection is performed

## Slide 28: Training Toolkit
Proprietary and confidential — do not distribute
Six Detailed Training Videos
PowerPoint Training
Helpful Tools & Support Documents
FAQs and Technical Service Contacts Monday-Friday 8am EST – 8pm EST 1-800-257-9525 or ts.scr@abbott.com
https://www.globalpointofcare.abbott/en/support/product-installation-training/navica-brand/navica-binaxnow-ag-training.html

## Slide 29: Additional Resources
Technical Support Line:
US +1 800 257 9525, Option 2
– 8am-8pm EST Monday-Friday
t s.scr@abbott.com

## Slide 30
The BinaxNOW™ COVID-19 Ag Test Card EUA has not been FDA cleared or approved. It has been authorized by the FDA under an emergency use authorization for use by authorized laboratories and patient care settings. The test has been authorized only for the detection of the presence of Antigen Protein from SARS-CoV-2, not for any other viruses or pathogens, and is only authorized for the duration of the declaration that circumstances exist justifying the authorization of emergency use of in vitro diagnostic tests for detection and/or diagnosis of COVID-19 under Section 564(b)(1) of the Act, 21 U.S.C. § 360bbb-3(b)(1), unless the authorization is terminated or revoked sooner.
© 2020 Abbott. All rights reserved. All trademarks referenced are trademarks of either the Abbott group of companies or their respective owners. Any photos displayed are for illustrative purposes only. 120006994-09 11/20

## Slide 31: Using the BinaxNOWTM COVID-19 Ag Cards
Kevin B. Spicer, MD, PhD, MPH, Medical Officer (CDC)
Antibiotic Resistance Coordinator, HAI/AR Prevention Program, KDPH

## Slide 32: Use-Cases for BinaxNOW Ag Test Cards
Three primary scenarios for effective use:
  Outbreaks
  Clinical settings
  “Surveillance” testing
    2020: Alternate with PCR
    2021: Medicare-Certified (from KDPH)
    Non-Medicare-Certified (by Provider)
Each requires different thought processes
  How often they will be used
  Who they will be used with
  What kind of follow-up is needed
TM

## Slide 33: Outbreak Scenarios
Facilities may want to partner with local health department
  To determine strategies for possible use
  Coordinate procurement of cards
  Assist with administration if needed
Test symptomatic people to get early assessment
  Speeds up isolation and quarantine decisions
May test asymptomatic if indicated during outbreaks
  High likelihood of exposure
  At least two days have passed since exposure
  May not have adequate supplies of cards to follow this strategy
  Not highly recommended
    Resource intensive
    Interpretation of results questionable

## Slide 34: Interpretation and F/U in Outbreak Settings
If symptomatic
  Tests positive: Initiate isolation and contact investigation
  Tests negative: F/U PCR test to confirm negative
    Isolate until results received
    If still negative, quarantine if indicated by situation, consider testing for other illnesses
If definite exposure but asymptomatic
  Should be quarantined to start with
  Positive result indicates infection – Isolate/contact trace - quarantine
  Negative result – continue with quarantine
If no exposure identified and asymptomatic
  Recommend confirming positive result with PCR
  Negative result taken at face value

## Slide 35: Testing of Symptomatic Persons in Clinical Settings
Can be used in healthcare settings for rapid assessment of COVID-19
  Must have CLIA certification or waiver:
  https://www.cms.gov/Regulations-and-Guidance/Legislation/CLIA/How_to_Apply_for_a_CLIA_Certificate_International_Laboratories
  FDA Emergency Use Authorization (EUA) approved for suspected COVID-19
  Should be used within the first 7 days of symptom onset
Consider Situation:
  If Yellow or Green on Indicator Map(or very few) cases identified in community (not likely currently!)
    Recommend F/U PCR test to confirm a positive test – Isolate until PCR results received
    Negative result taken at face value – may want to test for Flu, other respiratory pathogens
  If cases already prevalent in community
    Positive result would be taken at face value
    Negative results would be recommended for F/U PCR to confirm if there was a clear exposure UNLESS another etiology for symptoms is identified through testing
    Negative results do not eliminate the need for quarantine if exposed

## Slide 36: “Surveillance Testing” of Defined Populations
Goal is to prevent introduction into facility or patients
Targeted testing
  Recommended at least weekly (2X per week ideal)
  Staff AND In-coming Residents
If Asymptomatic (Employee or In-coming Residents)
  Positive results: recommend PCR confirmation (Isolate until results received)
  Negative results: taken at face value
If Symptomatic Employee – should NOT be at work!
  Get tested outside
  Isolate until results received
  Report results to workplace
If Symptomatic In-coming Resident
  Positive results: taken at face value (isolate and keep precautions in place)
  Negative results: confirmed with PCR
Must have strategy for follow-up PCR testing

## Slide 37: Reporting Requirements
Kevin B. Spicer, MD, PhD, MPH, Medical Officer (CDC)
Antibiotic Resistance Coordinator, HAI/AR Prevention Program, KDPH

## Slide 38: Reporting BinaxNow® Results
Interim solution pending access and reporting through NHSN or KHIE portal…
XLS Spreadsheet (flat file): https://chfs.ky.gov/agencies/os/oig/dhc/Pages/cvltc.aspx
Guidance for reporting antigen test results changed in January, available at:
https://chfs.ky.gov/agencies/dph/covid19/GuidancePOCAgtestingLTCFs.pdf

## Slide 39: Kentucky Health Information Exchange (KHIE) Portal https://khie.ky.gov/COVID-19/Pages/Direct-Lab.aspx
Charlese Blair, CRT, CMA
Project Coordinator, Office of Health Data and Analytics
Kentucky Health Information Exchange (KHIE)

## Slide 40: Requesting BinaxNOW Ag Test Cards
Complete and submit a Resource Request Form at: https://ky.readyop.com/fs/4iY5/ac9d
  By 12/23/20 for 2020 Delivery
  Be sure to complete all sections, to include justification
KDPH staff will review and process resource requests daily.
KDPH will directly distribute test cards from the state warehouse and/or through UPS
KDPH Points of Contact
James R. House, Master Exercise Practitioner
State Health Operations Center Community Testing Operations Section Chief
Work Cell: (502) 330-5950
Email: jamesr.house@ky.gov
Robbie Hume
State Health Operations Center Logistics Section Chief
Work Cell: (502) 892-8899
Email: robbie.hume@ky.gov
TM

## Slide 41: Q & A (12/14 & 12/16): Funding 1
If funding is restored, will PCR testing resume even in part for staff and residents?
  If further federal funding is announced, we can revisit.
  Nobody is REQUIRED to use only state-provided antigen test kits
  They are AVAILABLE for any Medicare-certified LTC provider to help meet the CMS-expected testing frequency.
  If you choose to go another route, there is just no state budget to reimburse for that.
Notes: New Slide 12/18
Some overlap/duplication, but addressing questions posed in the chat box from both sessions

## Slide 42: Q & A (12/14 & 12/16): Funding 2
Is there a cost to the facility for the BinaxNOW™ kits moving forward?
BinaxNOW™cards were federally allocated to the states at no cost.
The only new cost to providers will be PCR testing that confirms a positive antigen test (which your clinical lab can bill traditionally to the appropriate 3rd party).
Notes: New Slide 12/18

## Slide 43: Q & A (12/14 & 12/16): Funding 3
Is there going to be any incentive to pay an employee to input all the testing results into the reporting areas. It will take as much as 9 hours to input residents and staff?
We have over 300 staff, and 200 residents at my facility. Is the state going to offer assistance or guidance on how to get all these antigen tests done...especially if we are required to do twice a week testing?
Not at this time.
No – Provider controls how to best operationalize.
Notes: New Slide 12/18

## Slide 44: Q & A (12/14 & 12/16): Reporting 1
Just want to make sure Nursing homes will be reporting through NHSN for individual tests and then aggregate totals through KOG?
Medicare-certified facilities with Level 3 access for NHSN report through NHSN and do not report through KHIE.
Medicare-certified facilities without Level 3 access for NHSN and Non-Medicare certified facilities report through KHIE https://khie.ky.gov/COVID-19/Pages/Direct-Lab.aspx
ALL facilities:
  Report positives to the Local Health Department
  ALL positive and negative aggregate results must be reported daily to KDPH through the COVID-19 KY Aggregate Daily Lab Report system. Entities reporting aggregate results will need to register for access to submit data. The access request link can be found at: https://tinyurl.com/KyLabCovidAggRpt
(See slide 38)
Notes: New Slide 12/18

## Slide 45: Q & A (12/14 & 12/16): Reporting 2
I thought Nursing homes reported antigen testing the residents and staff through NHSN, not on this website?
If you are reporting testing results through NHSN, you should NOT report results through KHIE. That will result in duplicates in the system.
ALL positive and negative aggregate results are reported daily to KDPH through the COVID-19 KY Aggregate Daily Lab Report system.
Notes: New Slide 12/18

## Slide 46: Q & A (12/14 & 12/16): Reporting 3
We already report all POC antigen-testing results in KOG and NHSN for both the BinaxNOW™ and BD Veritor.
And now we are to do another report for the reporting on BinaxNOW™?
Continue to report all POC antigen testing through NHSN (Medicare-certified) or the KHIE portal (Non-Medicare-certified or Medicare-certified without access to NHSN reporting), report positives to the LHD, and report aggregate results to KDPH through the COVID-19 KY Aggregate Daily Lab Report system. Entities reporting aggregate results will need to register for access to submit data. The access request link can be found at: https://tinyurl.com/KyLabCovidAggRpt
Reporting of any point-of-care testing for SARS-CoV-2 performed by the facility, whether BinaxNOW™, BD Veritor, Quidel Sofia, etc., will be reported as per previous slide.
Notes: New Slide 12/18

## Slide 47: Q & A (12/14 & 12/16): Reporting 4
If facility has level 3 NHSN access, and they are reporting their POC results, is that taking the place of reporting through the KHIE portal? (We understood from a DPH representative that the NHSN was sufficient.)
Facilities reporting through NHSN do NOT have to report through KHIE.
Continue to report to NHSN, report positives to the LHD, and do aggregate reporting through the COVID-19 KY Aggregate Daily Lab Report system.
Whether using NHSN, the KHIE portal, or the spreadsheet to report the POC testing results, all results (both positive and negative) are to be reported.
Notes: New Slide 12/18

## Slide 48: Q & A (12/14 & 12/16): Reporting 5
If we report through KOG and NHSN, do we have to report through KHIE?
Apples & Oranges: KOG is needed to access both the KHIE portal and the Aggregate Daily Lab Report system.
Facilities reporting through NHSN do not have to do anything with KHIE. Continue to report to NHSN, report positives to the LHD, and submit aggregate reporting through COVID-19 KY Aggregate Daily Lab Report system.
ALL positive and negative aggregate results must be reported daily to KDPH through the COVID-19 KY Aggregate Daily Lab Report system. Entities reporting aggregate results will need to register for access to submit data. The access request link can be found at: https://tinyurl.com/KyLabCovidAggRpt
Notes: New Slide 12/18

## Slide 49: Q & A (12/14 & 12/16): Reporting 6
Spoke with the KDPH Lab surveillance team and was informed that since we (Nursing Home) have level 3 access through NHSN we report our POC to them and to GenTrack survey that will ask for aggregate numbers. Is that all that is required for nursing homes?
The only additional requirement is to report positive tests to the LHD.
Notes: New Slide 12/18

## Slide 50: Q & A (12/14 & 12/16): Reporting 7
Do we report to the KHIE portal if we are Medicare-certified and are still waiting on level 3 access for NHSN?
Can a CCRC report its AL results through its SNF (CCRC)?
If you are awaiting level 3 SAMS access for reporting through NHSN, then you can EITHER use the KHIE portal OR the spreadsheet at:
https://chfs.ky.gov/agencies/os/oig/dhc/Pages/cvltc.aspx under Reporting & Forms (instructions are on Tab 1)
Do not report using both.
This is a question for NHSN (NHSN@cdc.gov). In the interim, reporting should be made using either the KHIE portal or the reporting spreadsheet.
Notes: New Slide 12/18

## Slide 51: Q & A (12/14 & 12/16): Reporting 8
Would you please repost the aggregate reporting site link for Medicare Certified facilities again?
https://tinyurl.com/KyLabCovidAggRpt
Notes: New Slide 12/18

## Slide 52: Q & A (12/14 & 12/16): Reporting 9
For Non-Medicare-Certified Facilities: the KHIE portal is the only option to report results?
This demonstration was ~3 minutes for one single person (entry) so 100 employees weekly would be ~5 hours. Am I correct that this is what is being requested from the state?
Yes. Report all results to the KHIE portal and aggregate results to GenTrack (COVID-19 KY Aggregate Daily Report system)
Yes
Notes: New Slide 12/18

## Slide 53: Q & A (12/14 & 12/16): Reporting 10
So we are still doing the survey monkey? (The call on Monday indicated we do not need to complete the Survey Monkey anymore since we are doing GenTrack and KHIE)
Apples & Oranges: “Survey Monkey®” is a proprietary e-polling tool; the reference on Monday was re: a different survey instrument.
Continue using the KDPH/HAI’s Survey Monkey tool to report on COVID-19 cases at your facility among both residents and staff (new, recovered, active, ever positive, deaths).
Notes: New Slide 12/18

## Slide 54: Q & A (12/14 & 12/16): Reporting 11
Is the KHIF portal mandatory or if local DPH says use EPI form can be used is that acceptable?
Where do we report the numbers of negatives and positives daily?
EPID 200 form can provide information for the local health department, but does not substitute for laboratory reporting requirement when performing Ag testing onsite.
Submit aggregate reporting through COVID-19 KY Aggregate Daily Lab Report system.
Notes: New Slide 12/18

## Slide 55: Q & A (12/14 & 12/16): Reporting 12
We send specimens to our lab and we report all positives to the local health department…is that all we need to do?
Yes.
Laboratory reporting for specimens tested and resulted offsite is the responsibility of the laboratory.
The lab will also do the aggregate reporting.
You are to report positives to the local health department
Notes: New Slide 12/18

## Slide 56: Q & A (12/14 & 12/16): Reporting 13
Plz clarify: for reporting to the health department-"Human Infection 2019 Novel Coronavirus Person Under Investigation (PUI) and Case Report," still complete for all positives?
The reporting process desired by the local health department should be discussed with the local health department.
That reporting may require the fillable KY Case Report Form (https://chfs.ky.gov/agencies/dph/covid19/CDCPUIFormKYFillable.pdf) or the EPID 200 per the discretion of the LHD.
Notes: New Slide 12/18

## Slide 57: Q & A (12/14 & 12/16): Reporting 14
If we are private pay only do we have to do the KHIE reporting?
Once a patient is entered in the KHIE system, will future POC test results be entered through the observation tab?
Report through KHIE and submit aggregate reporting through COVID-19 KY Aggregate Daily Lab Report system.
No…if a manual data submitter has more than one result on the same patient, the manual data submitter can enter both results at the same time on the Observation Page. (This saves the submitter from entering all of the patient information twice.)
Notes: New Slide 12/18
No, that’s not the case. If a manual data submitter has more than one result on the same patient, the manual data submitter can enter both results at the same time on the Observation Page. This will save the submitter from entering all of the patient information twice.

## Slide 58: Q & A (12/14 & 12/16): Testing 1
CDC says trained health care provider can conduct the test; in Kentucky, what credentials are required to conduct the swab?
Does it have to be a nurse that completes the testing?
Testing personnel must have a high school diploma, at minimum, and be trained with periodic assessment on their ability to perform quality testing.
Due to health and safety concerns, it is recommended that all non-healthcare persons administer the test under the direction of a trained health care professional.
Notes: New Slide 12/18

## Slide 59: Q & A (12/14 & 12/16): Testing 2
None of our boxes came with a log sheet for recording results…?
There are no log sheets with the Kits.
Materials Provided Per Kit:
  Test Cards (40): A cardboard, book-shaped hinged test card containing the test strip
  Extraction Reagent (1): Bottle containing 10 mL of extraction reagent
  Nasal Swabs (40): Sterile swabs for use with BinaxNOW™ COVID-19 Ag Card test
  Positive Control Swab (1): Non-infectious recombinant SARS-CoV-2 nucleocapsid antigen dried onto a swab
  Negative Control Swab: The use of a sterile patient swab ensures appropriate negative results are obtained
  Product Insert (1)
  Procedure Card (1)
Notes: New Slide 12/18

## Slide 60: Q & A (12/14 & 12/16): Testing 3
Is a positive test mandated to be followed up with a PCR test?
If you have an asymptomatic positive Ag test, but a negative follow-up PCR, is an additional confirmation test required?
See: https://chfs.ky.gov/cv19/AntigenTestResultsGuidance.pdf
Not all positive antigen tests require a PCR follow-up test.
PCR result prevails.
Notes: New Slide 12/18

## Slide 61: Q & A (12/14 & 12/16): Testing 4
Does a Non-Medicare-Certified LTCF that continues the PCR testing bi-weekly have access to the free BinaxNOW™ cards?
If symptomatic staff or residents occur in between the PCR biweekly testing will PCR testing be covered in the interim?
Only through 12/30; this resource is intended for Medicare-Certified LTCFs after the CARES Act funding expires to help meet the CMS mandate for testing frequency.
Only through 12/30
Notes: New Slide 12/18

## Slide 62: Q & A (12/14 & 12/16): Testing 5
I have a ton of BinaxNOW™ cards; if I still have them available can I use until out?
Yes - first exhaust any unexpired inventory of antigen test kits that were previously provided by HHS, then initiate use of BinaxNOW™ COVID-19 (Ag) test cards available from KDPH.
Notes: New Slide 12/18

## Slide 63: Q & A (12/14 & 12/16): Testing 6
What are we using for the MRN for staff being tested?
Employee ID or
an alternate unique identifier
(Drivers License, Final 6 of SSN, etc)
Note: employee's HOME address
Notes: New Slide 12/18

## Slide 64: Q & A (12/14 & 12/16): Test Frequency 1
Once vaccines are administered, will the state continue to require surveillance testing for those that receive it, if they remain asymptomatic?
With your continued assistance and cooperation, this initiative can extend into the months ahead on a parallel track with the LTC vaccination program…in anticipation that the demand for surveillance testing will soon begin to decline.
Further guidance regarding testing frequency after vaccination is expected from CMS & CDC, but for now there will be no change and testing frequency for Medicare-certified LTCFs; others are recommended to test at least bi-weekly.
Notes: New Slide 12/18

## Slide 65: Q & A (12/14 & 12/16): Test Frequency 2
Can you please repeat how testing will be changing in January?
Currently AL centers are testing PCR with a lab 2x per week. Do we understand that we will be going to only 1x every other week with PCR lab in January?
Please refer to LTC Testing program guidance
https://chfs.ky.gov/cv19/LTCTestingProgramProviderGuidanceUpdate.pdf
Correct
Notes: New Slide 12/18

## Slide 66: Q & A (12/14 & 12/16): Contacts 1
What is the email address to send to if we do not have level 3 access through NHSN yet?
How do we get access to the KHIE portal to report these?
Where does a LTCF reporting aggregate results will need to register for access to submit data.
How do we order BinaxNOW™ cards?
  Further Info?
NHSN@cdc.gov
https://khie.ky.gov/COVID-19/Pages/Direct-Lab.aspx
https://tinyurl.com/KyLabCovidAggRpt
https://ky.readyop.com/fs/4iY5/ac9d thru 12/30/2020
BinaxNOW™: 1-800-257-9525 or ts.scr@abbott.com
  https://www.globalpointofcare.abbott/en/support/product-installation-training/navica-brand/navica-binaxnow-ag-training.html
Notes: New Slide 12/18

## Slide 67: Points of Contact
Logistics and Request Process
James R. House, MEP, State Health Operations Center Community Testing Operations Section Chief, KDPH
Work Cell: (502) 330-5950
Email: jamesr.house@ky.gov
Robbie Hume, State Health Operations Center Logistics Chief, KDPH
Work Cell: (502) 892-8899
Email: robbie.hume@ky.gov
Surveillance, Testing and Reporting
Kevin B. Spicer, MD, PhD, MPH, Medical Officer (CDC)
Antibiotic Resistance Coordinator, HAI/AR Prevention Program, KDPH
Phone: (502) 564-3261 x4468
Email: kevin.spicer@ky.gov
Doug Thoroughman, PhD, MS, State Epidemiologist (Acting), CDC Career Epidemiology Field Officer, KDPH
Phone: (502) 564-3418 x4315
Email: douglas.thoroughman@ky.gov
Abbott
Michelle Moran, Technical Consultant Abbott Rapid Diagnostics
Cell: (502) 693-8829
Email: michelle.moran@abbott.com
KHIE
Charlese Blair, CRT, CMA, Project Coordinator, Office of Health Data and Analytics, Kentucky Health Information Exchange (KHIE)
Email: Charlese.Blair@ky.gov
Office of the Secretary
Keith R. Knapp, PhD, MHA, HSE, CNHA
Sr Advisor on Adult Programs
Work Cell: 502-229-3184
Email: Keith.Knapp@ky.gov
Notes: 12/18 – added K2